Government of Pakistan

Drug Regulatory Authority of Pakistan

Ministry of National Health Services, Regulations & Coordination

***

MINUTES OF THE 2ND MEETING OF THE MEDICAL DEVICE BOARD (MDB)

HELD ON 27TH JUNE, 2016

2nd meeting of the Medical Device Board (MDB) was held in the committee room of TF Complex, G-9/4, Islamabad on 27th June, 2016. The meeting was chaired by Dr. Noor Muhammad Shah, Director, Medical Devices & Medicated Cosmetics, Drug Regulatory Authority of Pakistan.  The meeting was attended by the following:-

| S.No. | Name and Designation / Department | Position in the MDB |
| --- | --- | --- |
|  | Dr. Noor Muhammad Shah, Director Medical Devices & Medicated Cosmetics, DRAP, Islamabad. | Chairman |
|  | Mrs. Tehreem Sara, Additional Director (MD&MC), DRAP, Islamabad. | Ex officio Member/ Secretary MDB |
|  | Dr. M.Saleem Memon, Director CDD, Directorate General Health Services, Sindh, Hyderabad. (Nominee of Director General Health, Sindh) | Ex-officio Member |
|  | Mr. Akbar Jan, Chief Drug Inspector, Health Department, Khyber Pakhtunkhwa (Nominee of Director General Health, Khyber Pakhtunkhwa ) | Ex-officio Member |
|  | Brig. (R) Dilshad Ahmed Khan, Professor of Pathology & Director Research, National University of Medical Sciences, Rawalpindi. | Member |
|  | Prof. Dr. Umar Hayat, Professor and Head of Interventional Cardiology, Ayub Medical College, Abbottabad. | Member |
|  | Mr. Muhammad Sulman, Director, Certification Bodies, Pakistan National Accreditation Council (PNAC), Islamabad. | Co-Opted Member |
|  | Mr. Luqman Ali, System Analyst, Pakistan Institute of Medical Sciences, Islamabad. | Member |
|  | Dr. Abdur Rashid, Chairman Quality Control, Division of QA&LT, DRAP, Islamabad. | Co-Opted Member |
|  | Miss. Tazeen S. Bukhari, TRF Technical Consultant for Medical Devices, Technical Resource Facility, 16-BB, Defence Housing Authority, Phase-IV, Lahore. | Member |
|  | Mr. Asif Jalil, ADC (MD&MC), DRAP, Islamabad. | Co-Opted Member |

The meeting started with recitation of the Holy Quran and Darud on the Prophet (PBUH).

Item No.I.    Formulation of Guidance document on Good Distribution Practice for Medical 	          Devices (GDPMD).

The MDB in its 1st meeting discussed various elements of the Medical Devices Rules, 2015 and decided to formulate guidance documents on various procedures.  The MDB was of the view that the guidance documents needed at the earliest stage of implementation shall be made first.  The Board deliberated that the applicants for registration of medical devices have to obtain the relevant establishment licence i.e establishment licence for manufacturing or import. The establishment licence for import, as per rules, requires certificate of Good Distribution Practice for Medical Devices (GDPMD) from the registered/authorized CAB.  It was accordingly decided that guidance document for GDPMD shall be made in the first instance. For this purpose, the MDB constituted a committee consisting of Secretary MDB (Coordinator), Dr. Muhammad Nadeem Ahmad, Brig.(R) Dr. Waqar Azim Niaz, Mr. Muhammad Tahir Aziz and Assistant Drugs Controller (MD&MC) for drafting guidance document for GDPMD to be placed before the MDB for perusal and approval.

The aforesaid Committee accordingly drafted the guidance document for GDPMD (Annexure-A) for perusal and approval of MDB.

Decision:  The MDB after detailed deliberation formulated and approved the guidance document for GDPMD placed as Annexure-A.

Item No.II.     Guidance on preparation of Site Master File (SMF) to be prepared by importer/ distributor.

The Site Master File (SMF) is a document to be prepared by importer/ distributor containing concise Good Distribution Practice for Medical Devices (GDPMD) information about the storage, distribution, deliveries and secondary assembly operations carried out at the named site.

The draft format along with guidance for preparation of Site Master File to be prepared by importer/distributor formulated by the Division of Medical Devices & Medicated Cosmetics was placed before the Board for consideration.

Decision: The MDB after detailed deliberation approved the format for preparation of Site Master File (SMF) to be prepared by the importer/ distributor (Annexure-B).

Item No.III.    Amendments in Medical Devices Rules, 2015.

The MDB was informed that the Drug Regulatory Authority of Pakistan with approval of Federal Government has exempted from registration those Conformity Assessment Bodies which are accredited with Pakistan National Accreditation Council (PNAC) with required scope vide SRO.459(I)/2016.  Therefore, the exempted Conformity Assessment Bodies having the required scope now have become eligible to conduct conformity assessment of the establishments and their medical devices and issue concerned certificates including certificate of Good Distribution Practice for Medical Devices (GDPMD).

It was further informed that the Drug Regulatory Authority of Pakistan with approval of the Federal Government  vide SRO.317(I)/2016 has also extended the exemption period of Class D medical devices and their establishments from operation of Medical Device Rules, 2015 till 30th June, 2016.

Decision:     The MDB appreciated the actions and opined that the activities were timely acquired and are in line with the phase-wise implementation plan of the Authority agreed and noted the above statement.

Item No.IV.       Regulation of sale of Medical Devices.

It was told to the MDB that under Rule 69 (3) of Medical Devices Rules, 2015, “the Provincial Governments shall in accordance with section 6 of the Act regulate the sale of medical devices in the prescribed manner and may for that purpose make such orders, and issue such directions to the importers, manufacturers, stockists, retailers or other dealers of medical devices, as they may deem fit.”

After notification, copies of Medical Devices Rules, 2015 were sent to the Provincial Governments for further necessary action at their end.

Decision:  The MDB discussed the matter and opined that the provincial health departments may take necessary actions at their end under Rule 69 (3) of Medical Devices Rules, 2015.

Item No.V.   Change in the Structure of Committee (s).

While discussing the formulation of Guidance Documents it was informed that sometimes members cannot become available due to their official or personal engagements, the MDB may authorize the Chairman to Change in the structure of committee (s) constituted for formulation of Guidance Documents for various procedures so that the committee may dispose of the required matter timely.

Decision:  The MDB authorized its Chairman to change members in the Committee constituted for drafting of Guidance Documents for various procedures.

Meeting ended with the Slogan “Pakistan Zindabad”.

==========

Annexure-A

MDB/GD No.01

June, 2016

GUIDANCE ON GOOD DISTRIBUTION PRACTICE FOR

MEDICAL DEVICES (GDPMD)

[Rule 10, Medical Devices Rules, 2015]

Regulatory Requirements for Medical Device Safety and Performance

PREFACE

Distribution is an important activity in the integrated supply-chain of medical devices. Various personnel and entities are generally responsible for the product sourcing, procurement, transportation, delivery, storage, device tracking, installation, commissioning, service and maintenance and calibration need to be appropriately managed and controlled to ensure the safety and performance of medical devices at the point of use. In some cases, however, a person or entity is only involved in and responsible for certain elements of the distribution process. The level of risks associated with these activities may be of similar degree as those in the manufacturing environment and the lack of control over these activities may affect safety and performance of the devices. The Guidance on Good Distribution Practice for Medical Devices (GDPMD) is developed to elucidate the requirements for an appropriate management and control of these activities.

GDPMD specifies the requirements for a quality management system to be established, implemented and maintained by an establishment in carrying out activities in medical device supply-chain to comply with Pakistan medical device regulatory requirements as stipulated in Medical Device Rules, 2015 and DRAP Act, 2012. GDPMD requires an establishment to demonstrate its ability to maintain quality, safety and performance of medical devices in compliance with the Pakistan medical device regulatory requirement throughout the supply-chain.

The certification to GDPMD is to be conducted by the authorized conformity assessment body. The design and implementation of GDPMD by an establishment is dependent on the types, categories and classification of medical device, size and structure of the establishment and the processes employed.

It is the responsibility of the establishment to ensure that they are in compliance with all applicable laws. This document provides general guidance for GDPMD.  Although it has been tried that the accurate information be provided, however the Authority does not warrants its accuracy and completeness and therefore does not accept any liability for any error or omissions in this document or for any action or decision taken or not taken as a result of using of this document. In the event of any contradiction between the requirements of GDPMD and any written law, the latter shall prevail.

# CONTACT INFORMATION

For further information, please contact:

Medical Devices Division, Drug Regulatory Authority of Pakistan, TF Complex, 7 Mauve Area, Sector G-9/4, Islamabad.

Phone Nos. 051-9262180, 051-9262183.

Email: dir.md.mc@dra.gov.pk

Website: www.dra.gov.pk

Contents

1		PRELIMINARY

1.1		Objective

1.2		Scope and application

1.3		Definitions

2.       	QUALITY MANAGEMENT SYSTEM

2.1		General Requirement:

- 2.2				Control of Documents

- 2.3				Control of Record

- 2.4				Complaints

2.5       	Field Corrective Action (FCA) and Field Safety Notice (FSN).

2.6		Recall

2.7       	Return

2.8		Disposal of medical devices

2.9		Mandatory problem reporting

- 2.10     			Internal audit

2.11     	Management review

2.11.1		Review input

2.11.2		Review output

2.12		Corrective action

2.13		Preventive action

- 3		RESOURCE MANAGEMENT

3.1		Personnel

3.2		Training

- 3.3			Responsibility and Authority

- 3.4			Designated person

3.5		Premises and Facilities.

- 3.6			Cleanliness

- 3.7			Pest Control

- 3.8			Calibration.

4		AUTHORIZATION

- 5			SITE MASTER FILE.

6		STORAGE AND STOCK HANDLING

6.1		Receipt of Stock.

6.2		Storage Condition:

- 6.3			Stock Rotation

- 6.4			Delivery to Customers

6.5		Installation and servicing

7		TRACEABILITY

8		SEGREGATION OF COUNTERFEIT, ADULTERATED, MISBRANDED, SPURIOUS AND SUBSTANDARD MEDICAL DEVICES

9		OUTSOURCE ACTIVITIES

10		SECONDARY ASSEMBLY

10.1		General Requirements:

- 10.2			Assembly documents

- 10.3			Materials control

10.3.1		Medical devices to be re-packed.

10.3.2		Packaging materials

- 10.3.3		Medical device labeling.

- 10.4		Good assembly practices

10.5		Quality control

1. 	PRELIMINARY

1.1	Objective

The objective of this document is to ensure the quality, safety and performance of medical device during all aspects of medical device supply-chain, which include, but not limited to procurement, storage, transportation,  delivery, installation, commissioning, service and maintenance, calibration and after sales service; tracking, documentation and record-keeping practices.

1.2	Scope and application

1.2.1	This document is applicable to all parties involved in the supply-chain of medical device covering authorized representatives of foreign manufacturers, importers or distributors of medical devices in Pakistan.

1.2.2	The design and implementation of GDPMD by an establishment is dependent on the types, categories and classification of medical device, size and structure of the establishment and the processes employed. If any requirement in GDPMD is not applicable due to the type, category and classification of the medical device and supply-chain activities, a justification has to be provided for exclusion from fulfillment of that particular requirement.

1.3	Definitions

In this document, unless the context otherwise requires—

1.3.1	“adverse effect” means any debilitating, harmful or detrimental effect that the medical device has been found to have or to be likely to have on the body or health of humans or animals when such a medical device is used by or administered to humans;

1.3.2 	“adverse event” means any event or other occurrence that reveals any defect in any medical device or that concerns any adverse effect arising from the use thereof;

1.3.3	“Authority” means the Drug Regulatory Authority of Pakistan established under the DRAP Act, 2012;

1.3.4	“complaint” means any written, electronic or oral communication that alleges deficiencies related to the identity, quality, durability, reliability, safety or performance of a medical device that has been placed on the market;

1.3.5	“distribution” for the purpose of this document means all aspects of medical device supply-chain, which include, but not limited to procurement, storage, transportation,  delivery, installation, commissioning, service and maintenance, calibration and after sales services, tracking, documentation and record-keeping;

1.3.6	“distributor” for the purpose of this document means any establishment involved in distribution;

1.3.7	“field corrective action (FCA)” is an action taken by a manufacturer to reduce a risk of death or serious deterioration in the state of health associated with the use of a medical device, which may include—

- Return of a medical device to the manufacturer or its representative;

- medical device modification, which may include retrofit in accordance with the manufacturer's modification or design change, permanent or temporary changes to the labeling or instructions for use, software upgrades including those carried out by remote access,  modification to the clinical management of patients to address a risk of serious injury or death related specifically to the characteristics of the device;

- medical device exchange;

- medical device destruction; or

- medical advice given by manufacturer regarding the use of the device.

1.3.8	“field safety notice (FSN)” means a communication sent out by a manufacturer or its representative to the device users in relation to a FCA;

1.3.9	“premises” for the purpose of this document, means any location that is used for the activities dealing with medical devices, including storage, secondary assembly, etc.;

1.3.10	“packaging” for the purposes of this document, means the container and other packaging material in which the medical device is supplied;

1.3.11	“primary package” means element of packaging system that maintains the integrity and/or sterility of a medical device;

1.3.12	“regulatory requirement” means the requirement stipulated under DRAP Act, 2012 and Medical Devices Rules, 2015;

1.3.13	“secondary assembly” means the process of repackaging of a medical device from its original packaging into another packaging, without breach of the primary package, before the medical device is supplied.

# 2.         QUALITY MANAGEMENT SYSTEM

2.1	General Requirement:

The establishment shall establish, document, implement and maintain a quality management system in accordance to the requirements of GDPMD.

Where an establishment chooses to outsource any activity that may affect the quality of medical devices, the establishment shall ensure control over such processes.

The quality management system established should be sufficiently robust to meet external and internal factors.

# Guidance:

It is important to recognize that the structure and level of detail required in the    procedures should be tailored to the needs of the establishment.

The documentation required include:

- a site master file,

- documented procedures required by the GDPMD,

- documents needed by the establishment to ensure the effective planning, operation and control of its processes,

- records required by the GDPMD, and

- any other documentation specified by the regulatory authority.

All documented requirements, procedures and activities shall also be implemented and maintained.

# 2.2	Control of Documents

Documents required for GDPMD shall be controlled.

A documented procedure shall be established for the control of documents.

All documents shall be prepared, approved, signed and dated by an appropriate authorized person(s) and any change in person(s) permitted to carry out this task requires authorization.

Documents shall be reviewed regularly and kept up-to-date. If a document has been revised, a control system shall be established to prevent the unintended use of the superseded version.

# 2.3	Control of Record

Records shall be established and maintained to provide evidence of conformity to requirements of GDPMD. Record shall be legible, readily identifiable and retrievable.

A documented procedure shall be established to define the controls needed for the identification, storage, protection, retrieval, retention time and disposal of record.

The establishment shall retain the record for a period of time:-

- specified by relevant regulatory requirements, or

- at least equivalent to the lifetime of the medical device as defined by the product owner of the medical devices, whichever is the longest.

# Guidance:

All necessary record shall be kept including record for each purchase and sale, showing the date of purchase and supply, name of medical device, quantity received or supplied and name and address of supplier or consignee.  Record should ensure traceability of the origin and destination of products.

Where an electronic record system is used in place of a paper-based system, the system utilized should have built-in checks and balances to ensure the integrity of the record and to protect against unauthorized entries and the system should also incorporate audit trails for tracking changes.

# 2.4	Complaints

The establishment shall establish a documented procedure for handling of complaints regarding medical devices.

Any report of adverse event which requires regulatory reporting shall be reported to the regulatory authority as per procedure and regulatory requirements.

Records of the complaint, investigation and any subsequent actions taken shall be maintained.

# Guidance:

Procedure describing the actions to be taken in the handling of all written and oral complaints regarding a medical device should be available. There should be record for each individual complaint.

The procedure for handling complaints shall ensure that the complaints received are investigated and followed through, and that corrective actions are taken to prevent repeated complaints, and, where a decision is made to recall the medical device, the details of the recall.

Within the company, a person shall be designated to handle complaints. This person must have the authority to initiate investigations. All investigations should be documented in writing.

If a medical device defect is discovered or suspected in a batch, consideration should be given to determine whether other batches are also affected.

The investigation should take into consideration the condition and circumstances under which the medical device was distributed, stored and used.

Complaints regarding a medical device or its packaging distinct from those relating solely to matters within the control of the distribution chain, must be notified promptly to the product owner. When the complaint is an adverse event related to a medical device, the regulatory authority must be notified promptly as per rules.

An investigation report must be prepared with all corrective/preventative actions clearly stated.

Complaints records should be reviewed regularly for any indication of specific or recurring problems requiring attention.

2.5        Field Corrective Action (FCA) and Field Safety Notice (FSN).

The establishment shall—

  - establish documented procedures for handling of FCA and FSN;

  - define the responsibilities for planning, conducting and reporting of corrective actions in the documented procedure;

  - establish in writing a recall or withdrawal procedure in consultation with manufacturer;

  - inform the MDB prior to execution of FCA and FSN;

  - inform all customers to whom the medical device was distributed with the appropriate degree of urgency;

  - inform overseas counterparts on the FCA and FSN if the medical devices are exported;

  - request that the affected medical devices be removed immediately from usable stock and stored separately in a secure area until they are disposed of in accordance with manufacturers’ instructions; and

  - maintain records of all actions taken in connection with the FCA and FSN.

2.6	Recall

The establishment shall—

  - establish a documented procedure to effectively and promptly recall medical device known or suspected to be defective;

  - ensure that the system comply with the regulatory requirements;

  - the manufacturer and/or authorized representative shall be informed in the event of a recall;

  - where a recall is instituted by an entity other than the manufacturer and/or authorized representative, consultation with the manufacturer and/or authorized representative should, where possible, take place before the recall is instituted;

  - recall information shall be reported to the Authority; and

  - the progress of a recall process shall be recorded and a final report shall be issued which includes a reconciliation between delivered and recovered quantities of products.

2.7       Return

The establishment shall establish documented procedures for handling of returned medical devices.

All returned medical devices shall be segregated apart from saleable stock to prevent redistribution until a decision has been reached regarding their disposal.

The criteria for re-evaluation of the returned medical devices shall be documented. Records of the re-evaluation and any subsequent actions taken shall be maintained.

Guidance:

Any product returned to the establishment should be treated as a nonconforming product.

Written procedures describing the handling of returned medical devices and the corresponding records of all returns should be kept.

All returned medical devices should be kept separate from saleable stock to prevent redistribution until a decision has been reached regarding their disposal.

Some criteria for medical devices to be returned to saleable stock are:-

- the medical devices are in their original unopened containers and in good condition;

- it is known that the medical devices have been stored and handled under proper conditions;

- the remaining shelf life period is acceptable; and

- the medical devices have been examined and assessed by appropriate personnel. This assessment should take into account the nature of the medical device, any special storage conditions required, and the time that has elapsed since it was distributed. Special attention should be given to thermo-labile medical devices. Advice should be sought from the product owner as necessary.

The returned medical devices should only be formally released to saleable stock, following a satisfactory quality re-evaluation by a nominated, responsible person.

Medical devices returned to saleable stock should be placed in accordance with the stock rotation system established.

2.8	Disposal of medical devices

The establishment shall establish a documented procedure for the disposal of medical devices, keeping in view the regulatory requirements.

If the medical devices have not been immediately sent for disposal, they shall be kept in a clearly segregated area and identified so that they will not be sold inadvertently or contaminate other medical devices.

Records of the disposal shall be maintained.

# Guidance:

Medical devices that no longer meet specification (for example, expired, contaminated, damaged, defective, etc) should be disposed of properly keeping in view the regulatory requirements.

Control should be established to ensure that:-

- the status is clearly identified;

- the products cannot re-enter the distribution system, and;

- it is disposed of safely.

2.9	Mandatory problem reporting

The establishment shall establish documented procedure for incident/problem reporting to comply with the regulatory requirements, which include—

- the identification of the nature of the incident/problem;

  - the investigation;

  - the evaluation and analysis; and

  - the action to be taken.

Each incident report shall lead to a final report where corrective actions are applicable.

# 2.10      Internal audit

The establishment shall conduct internal audit at planned intervals to monitor the implementation of and compliance with the requirements of GDPMD.

The establishment shall define in a documented procedure, the responsibilities and requirements for planning and conducting audit and reporting of the results and maintenance of the audit record.

Actions to eliminate detected nonconformities and their causes shall  be taken without undue delay. Verification of the actions taken and the reporting of verification results shall be recorded.

# Guidance:

Internal audit should normally be conducted once a year.

Internal audits are conducted and recorded in order to monitor the implementation and conformance with requirements of GDPMD.

The result of the audits is usually stated in a written report indicating the nonconformities found. Timely action should be taken to eliminate the nonconformities and their causes. The audit results should be communicated to management for review.

The establishment should identify and implement any changes necessary to ensure and maintain the continued suitability and effectiveness of the quality system through the use of the quality policy, quality objectives, audit results, analysis of data, corrective and preventive actions and management review.

2.11      Management review

The top management shall review its quality management system at planned intervals, to ensure its continuing suitability, adequacy and effectiveness. This review shall include assessing opportunities for improvement and the need for changes to the quality management system.

Records from management reviews shall be maintained.

2.11.1		Review input

The input to management review shall include information on:

  - results of internal and external audits;

  - customer complaints/feedback;

  - GDPMD regulatory compliance system and medical device compliance;

  - surveillance and vigilance activities including field corrective actions, advisory notes, recalls and adverse event /incident reporting;

  - feedback from manufacturer;

  - feedback and directives from the Authority;

  - status of preventive and corrective actions;

- follow-up actions from previous management reviews;

- changes that could affect the GDPMD regulatory compliance system; and

- recommendations for compliance.

2.11.2		Review output

The output from the management review shall include any decision and action related to:-

- the corrective and preventive actions required;

- the effectiveness of the GDPMD regulatory compliance system and its compliance with the Pakistan medical device regulatory requirements; and

- resource needs.

# Guidance:

The top management should review the quality management system on a regular basis. An annual review is generally acceptable for an established and effective quality management system.

These include minutes of meeting which can be produced, distributed and stored on paper or electronically.

Records of the management review should include the identity of those taking part in the review and all points of the review including description of any corrective or preventive action to be taken. For any action determined, the responsibility for such actions, the resources, target dates for completion, etc should be identified.

2.12	Corrective action

The establishment shall—

  - take action to eliminate the cause of nonconformities in order to comply with GDPMD and regulatory requirements; and

  - a documented procedure shall be established to define requirements for:

    - reviewing nonconformities (including customer complaints);

    - determining the causes of nonconformities;

    - evaluating the need for action to ensure that nonconformities do not recur;

    - determining and implementing action needed, including, if appropriate, updating documentation,

    - recording of the results of any investigation and of action taken; and

    - reviewing the corrective action taken and its compliance with GDPMD and regulatory requirements.

2.13	Preventive action

The establishment shall—

- determine proactive action to eliminate the causes of potential nonconformities in order to comply with GDPMD and regulatory requirements and preventive actions shall be appropriate to the effects of the potential problems; and

- establish a documented procedure to define requirements for—

  - determining potential nonconformities and their causes;

  - evaluating the need for action to prevent occurrence of nonconformities;

  - determining and implementing action needed;

  - recording of the results of any investigations and of action taken; and

  - reviewing preventive action taken and its effectiveness.

# 3	RESOURCE MANAGEMENT

3.1	Personnel:

Personnel working at the premises shall be sufficient for complying GDPMD. Qualified person shall possess the prescribed qualification.

# 3.2	Training

The establishment shall

- provide training to key personnel to satisfy needs, and

- maintain training record.

# Guidance:

Training for personnel will be tailored to the person’s assignment. Typical training and education should cover the:

- nature of business,

- health, safety and environmental regulations and regulatory body related requirements,

- establishment’s policies,

- function of the personnel, and

- procedures and instructions of relevance to personnel.

Special training may be necessary for personnel dealing with certain categories of substances/materials such as chemicals, biological, radiation emitting or energy source components and products.

# 3.3	Responsibility and Authority

The establishment shall ensure that responsibilities and authorities are defined, documented and communicated within the establishment. The establishment shall establish the interrelation between all personnel who manage, perform and verify work that affects quality, and shall ensure the independence and authority to perform these tasks.

# Guidance:

The	establishment	may	document	job descriptions,	including	key responsibilities and authorities, in the Site Master File.

The personnel holding key responsibilities and authorities may also be documented as part of procedures.

# 3.4	Designated person

The establishment shall appoint designated person who, irrespective of other responsibilities, shall have the ultimate responsibility of:

- ensuring that processes needed for the quality management system are established, implemented and maintained,

- reporting to top management on the performance of the quality management system and any need for improvement, and

- ensuring the  promotion  of  awareness  of  regulatory  and  customer requirements throughout the establishment.

- Liaising with external parties on matters relating to the quality management system and regulatory requirements.

3.5	Premises and Facilities.

The establishment shall ensure that the premises and equipment used are suitable and adequate to ensure proper conservation and distribution of medical devices.  Fire extinguishers, smoke detector etc for control of fire shall also be a vital part.

# 3.6	Cleanliness

The establishment shall establish documented requirements for cleaning of premises, including frequency and methods.  Records of cleaning shall be maintained.

# Guidance:

Storage areas should be cleaned and accumulated waste removed at regular intervals. The frequency and methods of cleaning the premises and areas should be recorded.

No smoking, eating and drinking should be permitted in areas used for  storage and handling of medical devices.

# 3.7	Pest Control

The establishment shall establish a pest control programme to identify and prevent pest infestation.

Record of pest control programme shall be maintained.

# Guidance:

The storage area should be designed and equipped to prevent the entry of insects, rodents and other pests/animals.

There should be a pest control programme to identify and prevent pest infestation. Appropriate records should be kept for these purposes.

# 3.8	Calibration.

The instruments used for measuring and monitoring temperature and humidity shall be calibrated or verified for accuracy from accredited laboratories at specific intervals, or prior to use and the results of such calibrations or verifications shall be recorded and retained.

4	AUTHORIZATION

The establishment shall—

  - obtain appropriate authorization from the relevant party to become authorized representative, importer or distributor of medical devices; and

  - establish and maintain written agreement with the relevant party pertaining to supply of information required for regulatory matters relating to medical devices it deals with.

# 5	SITE MASTER FILE.

The establishment shall establish and maintain a Site Master File approved by the MDB and includes

- the scope of the GDPMD implemented, including the justification for any exclusion and/or non-application,

- the documented procedures established for the GDPMD, and

- information regarding the premises where activity is conducted.

6		STORAGE AND STOCK HANDLING

The establishment shall provide suitable and adequate storage to ensure proper conservation of the medical devices.  The establishment shall provide back up for electricity to ensure proper conservation of the medical devices during electric load shedding.

6.1	Receipt of Stock.

The establishment shall establish and implement inspection or other activities necessary to ensure that medical devices received meet specified requirements.

Medical devices subject to specific storage measures shall be   immediately identified and stored in accordance with the specified instruction(s).

Records of the verification shall be maintained.

Guidance:

The receiving bays should protect deliveries from bad weather during unloading.

Medical devices requiring special storage conditions (e.g. temperature and/or humidity or narcotics requiring additional security measures) should be placed in separate areas equipped to provide the desired conditions.

A list of such medical devices should be maintained and the medical devices properly identified.

# 6.2	Storage Condition:

Medical devices shall be stored under conditions specified to prevent deterioration by light, moisture, temperature or other conditions.

Storage conditions shall be monitored and recorded periodically, where appropriate.

Records of the storage conditions shall be maintained.

# Guidance:

The storage of medical devices should be carried out in buildings or parts of buildings that have been built for, or adapted to this purpose.

Buildings should protect medical devices from contamination and deterioration, including protection from excessive heat or undue exposure to direct sunlight.

Premises should be constructed, serviced and maintained regularly to protect stored medical devices, from all potentially harmful influences such as undue variations of temperature and humidity.

There should be adequate storage areas, and where applicable, physically separated zones for the orderly segregation of medical devices.

The storage areas should have adequate lighting and ventilation.

Medical devices should be stored off the ground and suitably spaced to permit cleaning and inspection. Pallets should be well maintained and kept in a good state of cleanliness.

There should be designated areas, wherever required, for quarantined, saleable stock, expired, rejected/damaged, recalled and returned medical devices. Alternative means of segregation should be considered if proven, to prevent mix-up.

Appropriate and suitable storage conditions should be provided for hazardous, sensitive and dangerous materials such as combustible liquids and solids, pressurized gases, highly toxic and radioactive substances.

Adequate precautions should be taken against spillage or breakage, attack by micro-organisms, contamination and cross-contamination.

Where controlled environmental storage conditions are required, these conditions should be continuously monitored and documented. Appropriate actions, on the premise, equipment and/or materials should be taken when the storage conditions are not met. As far as possible, the actual storage temperature should be expressed quantitatively. Where the storage temperature is not expressed quantitatively or stated (in terms of a range) on the labels of the registered medical device, Annex-1 of this document should be used as guidance.

Temperature of the storage areas must be measured at suitable predetermined intervals to show the maximum and minimum temperatures for the day and recorded. Real time temperature measurement should also be provided wherever required. Where applicable, humidity measurements should also be performed.

It is recommended that temperature monitors be located in areas that are most likely to show fluctuations.

# 6.3	Stock Rotation

The establishment shall establish a system to ensure stock rotation. Medical devices beyond their expiry date or shelf life shall be segregated from usable stock. They shall be clearly labelled as “Not for Sale” or in other similar phrases/words.

The expired medical devices shall be disposed of as per procedure in this guidance document.

# Guidance:

Comprehensive records should be maintained showing all receipts and issues of medical devices according to batch number or equivalent.

Periodic stock reconciliation should be performed comparing the actual and recorded medical devices quantity. In any case, this should be performed when each batch is totally exhausted. All significant stock discrepancies should be subjected to investigation to check for inadvertent mix-ups and wrong issues.

A system should be in place to ensure that medical devices due for expiry earliest are sold and/or distributed first (Earliest-Expiry-First Out, EEFO). For cases where no expiry dates exist for the medical devices, FIFO (First-In- First-Out) should be applied.

Medical devices with broken seals, damaged packaging or suspected tampering/contamination must not be sold or supplied.

Medical devices bearing an expiry date must not be received or supplied  close to or after the expiry date such that this date is likely to occur before the consumer uses the medical devices.

All labels and containers of medical devices should not be altered, tampered or changed.

# 6.4	Delivery to Customers

The establishment shall establish adequate methods of transportation to achieve safe and secure delivery of all medical devices from their point of collection to their point of delivery.

Medical devices shall be transported in such a way that:-

- their identification is not lost;

- they do not contaminate, and are not contaminated by, other medical devices or materials/substances;

- adequate precautions are taken against spillage, breakage or theft;

- they are secure and not subjected to unacceptable degrees of heat, cold, light, moisture or other adverse influence, or to attack by microorganisms and pests.

Medical devices requiring controlled temperature storage or other special control and conditions shall be transported by appropriate or specialized means.

# Guidance:

Procedures should be in place to achieve safe and secure delivery of all medical devices from their point of collection to their point of delivery.

There should be adequate provision for the security, storage condition and protection of the quality of medical devices during all transportation. The transport process should not affect the integrity and quality of the medical devices.

Medical devices should be supplied to an authorized person only. When the    medical device is intended solely for professional use, such medical device should be  supplied to the targeted type of user.

Medical devices requiring controlled temperature storage should be transported by appropriate or specialized means and monitoring of temperature during delivery is recommended. Such records should be reviewed. Special care should be exercised when using dry ice during transportation. Medical devices should not come into contact with dry ice as this may cause freezing of the medical devices.

Vehicles used should be adapted and maintained to suit the operations to be carried out. Vehicles should not be used as a store for medical devices.

Deliveries should be made only to authorized wholesalers, distributors or persons authorized to supply medical device products in accordance with DRAP Act, 2012 and Medical Devices Rules, 2015.

6.5	Installation and servicing

The establishment shall establish and maintain documented procedures and work instructions for performing installation, testing and commissioning and maintenance activities in accordance with international standards and any other requirement as specified by the Authority.

The establishment shall—

  - establish and maintain documented procedures, work instructions and reference materials, tools and test equipment and reference measurement procedures, for performing servicing activities including calibration, repair, maintenance and verifying that they meet the regulatory requirements and applicable standards;

  - establish documented requirements which contain acceptance criteria for installation, testing and commissioning of the medical device;

  - establish installation qualification and maintain adequate installation and inspection instructions for medical devices requiring specified installation requirements, and where appropriate, test procedures;

  - ensure proper installation, testing and commissioning;

  - ensure equipment used for testing, maintenance and conservation of medical devices are calibrated or verified at specific intervals;

  - ensure the calibration and maintenance of test equipment conforms to the applicable standards; and

  - maintain testing and commissioning, installation, calibration and maintenance service records.

The establishment shall, as appropriate—

- establish an appropriate technical support which include maintenance service, training, calibration, management of spare parts, workshop setup and management;

- establish maintenance management mechanism to support the customers;

- ensure the technical and maintenance support services for active medical devices conform to the applicable regulatory requirements.

7	TRACEABILITY

Records providing traceability of medical devices from supplier and to the customers shall be maintained.

The establishment shall retain the record for a period of time:-

- specified by relevant regulatory requirements, or

- at least equivalent to the lifetime of the medical device as defined by the product owner of the medical devices, whichever is the longest.

# Guidance:

Identification of product by batch/lot/serial number permits traceability in two directions: forward to the customers and backward to the manufacturer.

Traceability to the customers is important as it is necessary to track medical devices right down to the user (e.g. Patients or hospitals) in the event of a  field safety corrective action (FSCA). The traceability to the manufacturers allows for investigation of quality-related problems and feedback to prevent recurrence of nonconforming products.

Records providing traceability may include delivery order from the supplier, distribution or sales records etc. The records should allow for the date, name of medical device, quantity supplied, lot/batch number or serial number, name and address of the supplier and addressee to be ascertained.

Proof of delivery transactions should also be kept.

8.	SEGREGATION OF COUNTERFEIT, ADULTERATED, MISBRANDED, SPURIOUS AND SUBSTANDARD MEDICAL DEVICES

# Counterfeit, adulterated, misbranded, spurious and substandard medical devices, if found in the distribution network, shall be physically segregated from other medical devices to avoid any confusion. They shall be clearly labelled as “Not for Sale” or in other similar phrases/words.

The establishment shall inform the regulatory authority and manufacturer immediately.

# Guidance

Counterfeit, adulterate, misbranded, spurious and substandard medical devices are defective products that require reporting by the establishment to the regulatory authority. Any possession of aforesaid medical devices by the establishment must be reported to the regulatory authority.

Reporting shall be performed as per procedure. Once the establishment becomes aware of possession of possibly counterfeit, adulterated, misbranded, spurious and substandard medical devices should be reported immediately. Key information (i.e. brand, place and period of purchase) should be provided in the report.

Appropriate corrective measures should be undertaken for identified counterfeit, adulterate, misbranded, spurious or substandard  medical device, which includes but are not limited to the following:-

- Reporting of possession of these medical devices to the regulatory authority,

- Segregation/quarantine of these medical devices,

- Investigation of supply chain breach,

Communication to all affected wholesalers / distributors / retailers / consignees possibly supplied with these medical devices (prior notification to the regulatory authority is required for such communication

9.      OUTSOURCE ACTIVITIES

Where the establishment outsources any process within the scope of the GDPMD, the establishment shall ensure control over such processes.

The establishment shall establish requirements to ensure that the outsourced activities conform to specified requirements.

The type and extent of control applied to the supplier shall be dependent on the impact on meeting the requirements of GDPMD.

The supplier shall be audited as part of the internal audit if the supplier has not been certified to GDPMD.

For outsourced activities, the supplier of out- sourced storage, warehousing, secondary assembly and distribution services shall be audited by the certification body as part of the establishment’s system, unless the supplier is already certified to GDPMD. The scope of the supplier's certification shall cover the scope of the establishment.

# Guidance:

Certain processes such as cleaning, pest control, transportation, storage etc may be outsourced to third party service providers.

The establishment is required to demonstrate that it has control over the outsourced process by maintaining evidence that the selection of the supplier was based on assessment appropriate to the service being purchased and the ability of the supplier to enable the establishment to meet requirements of GDPMD and any other regulatory requirements.  When a supplier carries out any operation, a written contract specifying the limits of responsibility of the establishment and the supplier should be considered.

When monitoring the performance of the suppliers, the establishment should consider compliance trends and conformance history. The performance of the supplier should be monitored on a regular basis. Depending on the nature of the activities performed, the service provider may be subjected to inspection by the regulatory authority.

10.	SECONDARY ASSEMBLY

10.1	General Requirements:

For the purpose of secondary assembly, the establishment shall plan and carry out secondary assembly of medical devices under controlled conditions.

Controlled conditions shall include, as applicable:

- the availability of information that describes the characteristics of the medical device,

- the availability of documented procedures, documented requirements, work instructions, and reference materials and reference measurement procedures as necessary,

- the use of suitable equipment,

- the availability and use of monitoring and measuring devices,

- the implementation of monitoring and measurement activities,

- the implementation of release of medical devices, their delivery and post-delivery activities, and

- the implementation of defined operations for labelling and packaging of medical devices.

The establishment shall establish and maintain a record for each batch of medical devices that provides traceability for a period mentioned in this document and identifies the amount assembled and the amount approved for distribution. The batch record shall be verified and approved by qualified personnel.

Guidance:

The size of assembly area must reflect the volume of assembly. The adequacy of the working space should permit the orderly and logical positioning of equipment and materials so as to avoid confusion and to minimize the risk of mix-up between different medical devices or their components.

Assembly areas should be well-lit and effectively ventilated, with air handling system (including temperature, humidity and filtration controls) where appropriate.

No smoking, eating and drinking should be permitted in areas used for storage and handling of medical devices.

Segregated areas should be provided for the storage of approved, quarantined, rejected, recalled and returned materials or products.

In considering which controlled conditions are applicable for a given process, a due consideration an establishment should make is if in the absence of the control, whether there could be an adverse or potentially adverse affect on quality of the medical device. If yes, then control is necessary. The amount of control and level of detail may be commensurate with the degree of   criticality (e.g. based on the output of risk management activities) of the process in achieving the requirements for quality and the degree of training of personnel involved in the secondary assembly line.

Reference materials may be physical or visual, such as product examples indicating permissible colour variation, or visual such as photos of known non-conformities. Reference materials should be available at the point of use. An individual procedure could be in the form of a simple flowchart, or a processing sequence, combined with a checklist.

Suitable equipment should be designed and selected so that process and product specifications are met. It should be verified that new and/or significantly modified equipment meets purchasing/design specifications and is capable of operating within its defined limits and the process operating limits.

The risk of labelling and packaging errors can be minimized by the introduction of appropriate controls such as:-

- segregation	of	secondary	assembly	(e.g.	packaging	and labelling) operations from other operations,

- avoidance of packaging and labelling product of similar appearance in close proximity,

- secondary assembly line identification,

- application of line clearance procedures,

- destruction of unused batch-coded materials on completion of packaging and labelling,

- use of roll-feed labels,

- use of a known number of labels and reconciliation of usage,

- on-line printing, including batch coding,

- use of electronic code encoders/readers and label counters,

- use of labels designed to provide clear product differentiation,

- inspection of label content before use, and

- proper storage of labels in areas of restricted access.

Any secondary assembly work instructions should include:

- Name of product;

- Description of the applicable medical devices for assembly and pack size;

- A complete list of all the packaging materials required for a standard batch size, including quantities, sizes and types, with the code or reference number relating to the specifications of each packaging material;

- Where appropriate, an example or reproduction of the relevant printed packaging materials, and specimens indicating where to apply batch number references, and shelf-life of the product;

- Special precautions to be observed, e.g. a careful examination of the area and equipment in order to ascertain the line clearance before operations begin;

- A description of the packaging operation, including any significant subsidiary operations, and equipment to be used;

- Details of any in-process controls with instructions for sampling and acceptance limits, if applicable.

# 10.2	Assembly documents

A batch assembly record shall be kept for each batch or part batch assembled. The record shall carry the batch number and the quantity of bulk medical devices to be packed.

The assembly shall be made or completed at the time each action is taken and in such a way that all significant activities concerning the assembly of medical device are traceable.

These records shall be retained for a period of time specified by relevant regulatory requirements, or at least equivalent to the lifetime of the medical device as defined by the product owner of the medical devices, whichever is the longest.

# Guidance:

Records that facilitate traceability and review of the secondary assembly of a batch of product, derived during the secondary assembly of that batch, should be contained in a batch record, and are frequently collated in a single file. Such files can be referred to as a “Device History Record”, “Batch Assembly Record”, “Lot History Record” or “Lot Record”

If it is not practical to include all the relevant documents in the batch record, then the record should list the titles of those documents and their location(s).

During secondary assembly, relevant information should be entered onto the batch record. Such information can include:

- the quantity of raw materials, components and intermediate products, and their batch number, if appropriate,

- the date of start and completion of different stages of secondary assembly, including sterilization records if appropriate,

- the quantity of product assembled,

- the signed results of all inspections and tests,

- the designation of the product line used, and

- any deviation from the secondary assembly specifications.

Data may be recorded by electronic data processing systems, photographic or other reliable means, but detailed procedures relating to the system in use should be available and the accuracy of the records should be checked.

If documentation is handled by electronic data processing methods, only authorized persons should be able to enter or modify data in the computer and there should be a record of changes and deletion; access should be restricted by passwords or other means and the result of entry of critical data should be independently checked.

Batch records electronically stored should be protected by back-up transfer on magnetic tape, microfilm, paper or other means. It is particularly important that the data are readily available throughout the period of retention.

Records should be maintained of the distribution of each batch of the product in order to facilitate the recall of the batch if necessary.

# 10.3	Materials control

10.3.1	Medical devices to be re-packed.

For  each  delivery,  the  incoming  medical  devices  shall  be  checked  for integrity of package and seal, for correspondence between the delivery note and the supplier’s labels, and for compliance with medical device quality specification.

Medical  devices  with  breached  primary  package  shall  not  be  used  for secondary assembly.

Medical devices in the storage area shall be appropriately labelled. There shall be appropriate procedures or measures to assure the identity of the contents of each packaging of the medical devices. Bulk containers from which quantities of the medical devices have been drawn shall be clearly identified.

Medical devices requiring special storage conditions shall be placed in separate areas constructed and equipped to provide the desired conditions. The storage conditions shall be continuously monitored and recorded.

As far as possible, the actual storage temperature shall be expressed quantitatively. Where the storage temperature is not expressed quantitatively or stated (in terms of a range) on the labels of the registered medical device, the guidance at Annex-1 shall apply.

10.3.2		Packaging materials

The purchase, handling and control of all packaging materials shall be accorded attention similar to that given to starting materials.

When setting up a programme for the packaging operations, particular attention shall be given to minimizing the risk of mix-ups or substitutions. Different medical devices shall not be packaged in close proximity unless there is physical segregation.

# 10.3.3	Medical device labeling.

All original labelling (including instructions for use, label and any other informational sheet or leaflet, etc) and all original labelling information shall accompany the repackaged medical device when it is supplied alongwith secondary package labelling.

# Guidance:

Medical devices should only be purchased from approved suppliers named in the relevant specifications. A batch certificate of analysis should be obtained from the supplier for each batch of product to be re-packed.

Inspection of breached primary packages has to be performed. Any breached packages should be disposed off and should not be re-processed and returned to the secondary assembly line.

Manufacturing activities such as sterilization do not fall within the definition of secondary assembly.

The Quality Management System (QMS) Requirements for all other manufacturing activities other than secondary assembly is ISO 13485 Medical devices – Quality management systems – Requirements for regulatory purposes.

Particular attention should be paid to printed materials. They should be stored in adequately secure conditions such as to exclude unauthorized access. Cut labels and other loose printed materials should be stored and transported in separate closed containers so as to avoid mix-ups. Packaging materials should be issued for use only by authorized personnel following an approved and documented procedure.

Each delivery or batch of printed or primary packaging materials should be given a specific reference number or identification mark.

Outdated or obsolete primary packaging material or printed packaging material should be destroyed and the disposal recorded.

# 10.4	Good assembly practices

All medical devices and materials used for assembly shall  be checked before use by a designated person for quantity, identity and conformity with the packaging instructions. Line clearance shall be performed prior to commencement of the assembly operation.

The  correct  performance  of  any  printing  operation  which  is  carried out separately or in the course of the packaging shall be checked and recorded.

The establishment shall ensure that  the  assembly  equipment  used    are suitable and adequate to ensure proper secondary assembly of medical devices.

The parts that come into contact with the medical devices must not affect the quality of the medical devices and present any hazard.

Control equipment shall be calibrated and checked at defined intervals by appropriate methods. Adequate records of such tests shall be maintained.

Guidance:

Before any assembly begins, there should be recorded checks that the equipment and work station are clear of previous products, documents or materials not required for the planned process, and that equipment is clean and suitable for use.

Each batch of products produced must be assigned a unique batch number. Normally, filling and sealing should be followed as quickly as possible by labeling. If it is not the case, appropriate procedures should be applied to ensure that no mix-ups or mislabeling can occur.

The performance of the printing operation should be checked and recorded. Attention should be paid to printing by hand which should be re-checked at regular intervals.

Printed and embossed information on packaging materials should be distinct and resistant to fading or erasing.

On-line control of the product during packaging should include at least checking the following:

- General appearance of packages;

- Completeness of packages;

- Correctness of products and packaging materials used;

- Correctness of over-printing, e.g. batch number, expiry date.

Any significant or unusual discrepancy observed during reconciliation of the amount of bulk product and printed packaging materials and the number of units produced should be investigated and satisfactorily accounted for before release.

Upon completion of a packaging operation, any unused batch-coded packaging materials should be destroyed and the destruction recorded. A documented procedure should be followed if un-coded printed materials are returned to stock.

Where any packaging process validation applies, the study should be carried out in accordance with pre-defined and authorized protocols.

The appropriate marketing authorization of the product(s) to be assembled shall be complied with, and where necessary, the approval of the holder of the market authorizations and licensing authority shall be sought.

10.5	Quality control

Finished medical device assessment shall embrace all relevant factors, including assembly conditions, a review of packaging documentation, compliance with finished medical device specification and visual examination of the final finished pack.

The process of secondary assembly shall not compromise the medical device’s conformance to the Essential Principles of Safety and Performance.

# Guidance:

There should be designated and independent personnel in charge of Quality Control (QC).

If sampling is performed, the samples taken should be done in accordance with approved written procedures that describe:

- The method of sampling;

- The equipment to be used;

- The amount of the samples to be taken;

- The storage conditions for the samples taken

Any in-process controls, including those carried out in the assembly area, should be approved by QC and the results recorded.

# ANNEX 1

T1: Storage Temperature Values

| ON THE LABEL | GUIDANCE VALUES |
| --- | --- |
| Freezer | The temperature is thermostatically controlled between –20˚C and –10˚C |
| Refrigerator | The temperature is thermostatically controlled between 2°C and 8°C. |
| Cold place | The temperature does not exceed 8°C. |
| Cool place | The temperature is between 8°C and 15°C |
| Room temperature | The temperature is between 15°C and 30°C |
| Warm | The temperature is between 30°C and 40°C. |
| Excessive heat | The temperature is above 40°C. |
| Do not store over 30°C | The temperature is between 2°C and 30°C. |
| Do not store over 25°C | The temperature is between 2°C and 25°C |
| Do not store over 15°C | The temperature is between 2°C and 15°C |
| Do not store over 8°C | The temperature is between 2°C and 8°C. |
| Do not store below 8°C | The temperature is between 8°C and 25°C. |

T2: Storage Condition Guidelines

| ON THE LABEL | GUIDANCE VALUES |
| --- | --- |
| Protect from moisture | No more than 60% relative humidity in normal storage conditions; to be provided to the user in a moisture- resistant container. |
| Protect from light | To be provided to the user in a light resistant container. |

ANNEX- 2

SCOPE OF GDPMD CERTIFICATION

- The GDPMD certification shall specify the following—

    - scope of activities performed by the establishment and medical devices dealt with by the establishment;

    - outsourced activities, if applicable;

    - any special storage and handling conditions, such as chill room or cold room for cold chain management; and

    - applicable sections of the Medical Device Rules, 2015 and the relevant laws.

- The following information shall be indicated in the certification of GDPMD—

    - name, complete address and contact information of the establishment involved in performing activities 	within the scope of GDPMD certification and the activities performed including outsourced 	activities within the scope of GDPMD certification.

  - particulars of CAB issuing the certificate which include the name and address, logo, registration number where applicable and/or accreditation number issued by PNAC and the name and signature of the certification manager of the CAB; and

  - number, issuing date and expiry date of the certificate.

- Scope of activities for establishment including those activities that are outsourced to be certified include anyone or combination of the following activities—

    - import;

    - storage and warehousing;

    - secondary assembly;

    - distribution including transportation;

    - installation, testing & commissioning; and

    - servicing and maintenance; and

(4) 	List of devices dealt with by the establishment—

| S.No | *Device Category |
| --- | --- |

*List of device categories:

01 Active implantable devices

02 Anesthetic and respiratory devices

03 Dental Devices

04 Electro mechanical medical devices

05 Hospital hardware

06 In vitro diagnostic devices

07 Non-active implantable devices

08 Ophthalmic and optical devices

09 Reusable devices

10 Single-use devices

11 Assistive products for persons with disability

12 Diagnostic and therapeutic radiation devices

13 Complementary therapy devices

14 Biologically-derived devices

15 Healthcare facility products and adaptations

16 Laboratory equipment

17 Medical software

18 Others: Please specify with justification for any additional categories

Annexure-B

SITE MASTER FILE

| S.No. | Requirements | Guidance |
| --- | --- | --- |
| 1 | GENERAL INFORMATION |  |
| 1.1 | Name and exact address of the site, including telephone and fax numbers | - Name of company, site address and mailing address (if different from site address) - Telephone, fax nos. and email address of contact person |
| 1.2 | Brief information on the site. | Outline the company's activities. This covers both medical device related and non-medical device related activities. Secondary assembly activities, if performed, are to be specified. |
| 1.3 | Short description of the site (location and area of site) | - Provide a map indicating the location of the site. - Area of site |
| 1.4 | GDPMD operations as licensed by the Competent Authorities | - Indicate whether the premises has obtained approval from any authority in Pakistan, or any foreign Competent Authority - Quote the relevant document (licence) as issued by the Competent Authority. State the period of validity of licence document. |
| 1.5 | Number of employees engaged in administration, warehousing, distribution, transportation, secondary assembly and others. | - Administration - Warehousing - Distribution - Transport - Secondary assembly (if applicable) - Others (if any) |
| 1.6 | Class(es) of medical devices handled on the site | Mention the Class(es) of medical devices handled. |
| 1.7 | Information about handling of medical devices and other specifically toxic or hazardous substances handled and mention the way they are handled and precautions to be taken. | Mention medical devices needing special storage conditions and any toxic, hazardous or other substances needing special storage conditions. Mention the way they are handled and precautions to be taken. |
| 1.8 | Use of outside administrative or other technical assistance in relation to the operation. | For each work process outsourced or sub-contracted (including contract delivery companies), give:- - Name, address, telephone no. and fax. no. of contractor - Brief outline of the activity being undertaken. |
| 2 | PERSONNEL |  |
| 2.1 | Organization chart showing the arrangements for key personnel | Organogram listing key personnel. Record senior managers and supervisors only. |
| 2.2 | Qualifications, experience and responsibilities of key personnel | - Brief details of qualifications and relevant experience. - Job descriptions for the key personnel |
| 3 | PREMISES AND FACILITIES |  |
| 3.1 | Simple layout plan and description of warehousing areas | - Layout of premises. - Provide a site layout plan highlighting all warehousing and other functional areas. - Describe the controls available to prevent unauthorized access. |
| 3.2 | Brief description of ventilation systems. More details should be given for critical areas providing special storage conditions | Brief description of ventilation systems etc. More details should be given for critical areas. To reduce the narrative, schematic drawings may be used. |
| 3.3 | Procedures for cleaning the areas | - Briefly describe cleaning procedures for the areas and the frequency? |
| 3.4 | Policy on the storage of medical device | - How are medical devices of different status e.g quarantined, rejected, approved, etc segregated and controlled, e.g. labels? - How are the medical devices stored e.g. pallet racking? - Describe any special storage or handling conditions such as cold chain management. - Describe the pest control programme. |
| 4 | DOCUMENTATION |  |
| 4.1 | Documentation used in GDPMD activities | This section refers to documentation used in GDPMD activities. - Who is responsible for the preparation, revision and distribution of documents? - Where are the master documents stored? - Is there a standard format and instruction of how documents are to be prepared? - How is the documentation controlled? - For how long are the documents kept? - any arrangement for electronic records. |
| 5 | CONTRACT ACTIVITIES (if any) |  |
| 5.1 | Description of the contract. | Describe briefly the details of the contract between the contract giver and acceptor. The type of activities undertaken by the contract acceptor should be specified. |
| 6 | SECONDARY ASSEMBLY (if applicable) |  |
| 6.1 | Description of the secondary assembly process | - Medical devices that are to be repackaged, the equipment used in the assembly activities. - List of applicable work instructions and operational procedures for secondary assembly. |
| 6.2 | Description of the documentation process for traceability | - Description of the documentation system for batch assembly records? - Who is responsible for verifying and approving the batch assembly records? - Is there a standard format and instruction of how documents are to be prepared? - How is the documentation controlled? - For how long are the batch assembly records retained? - Does the batch assembly record identify the amount assembled and approved for distribution? |
| 6.3 | Description of the materials control established | - Prior to repackaging, - Are the incoming medical devices checked for integrity of package and seal? - Are the incoming medical devices appropriately labeled? - (If applicable) are the special storage conditions for the incoming medical devices adhered to? - Are the packaging materials appropriately purchased, handled and controlled? - Do the repackaged medical devices bear all the original labeling information (including instructions for use, label and any other informational sheet or leaflet, etc.)? |
| 6.4 | Description of the good assembly practice | Written procedures for the following, if performed:- - Assembly operation - Printing operations - Cleaning and maintenance of assembly equipment - Calibration of assembly equipment |
| 6.5 | Description of quality control implemented | Written procedures for finished medical device assessment, including:- - Review of packaging documentation, - Compliance with finished medical device specification and examination of finished product. |

****************